## Slide 1: Overview: US FDA Proposed Rule Quality Management System Regulation
Melissa Torres
Associate Director for International Affairs
Office of the Center Director
Center for Devices and Radiological Health
U.S. Food and Drug Administration

## Slide 2: Overview
Discuss the proposed amendment to 21 CFR Part 820: Quality Management System Regulation (QMSR)
Review ISO 13485:2016 and understand the rationale for its incorporation by reference into the proposed rule
Identify the changes in the proposed QMSR

## Slide 3: Proposed Quality Management System Regulation (QMSR)
FDA published the proposed amendment to 21 CFR Part 820: Quality Management System Regulation (QMSR) on February 24, 2022
https://www.federalregister.gov/documents/2022/02/23/2022-03227/medical-devices-quality-system-regulation-amendments
Revisions to Part 820 replace most of the existing regulation with an incorporation by reference (IBR) to the 2016 edition of International Organization for Standardization (ISO) 13485 - Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes
90-day comment period closed on May 24, 2022

## Slide 4: ISO 13485: 2016 Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes
Establishes requirements for a Quality Management System (QMS) that can be used by an organization involved in one or more stages of the life-cycle of a medical device, including:
  design and development
  production
  storage
  distribution
  installation
  servicing and
  final decommissioning/disposal of medical devices

## Slide 5: Rationale for Utilization of ISO 13485: 2016
Regulatory expectations for a QMS have evolved
Used by many other countries
Lessons learned through current and previous FDA programs utilizing ISO 13485 demonstrated feasibility
Requirements are substantively similar between the current part 820 and ISO 13485:2016

## Slide 6: Benefits of ISO 13485: 2016
  Modernized QMS principles
  Greater integration of risk management activities
  Globally harmonized requirements
    Standard used by many other Regulatory Authorities
    Many global manufacturers already comply with ISO 13485

## Slide 7: Goals of the Proposed Quality Management System Regulation
Simplify and Streamline the Regulation
Reduce burden on manufacturers
Keep country specific requirements at a minimum
Maintain a similar level of assurance in a manufacturer’s quality management system

## Slide 8: Overview of the Proposed Quality Management System Regulation
Withdraws most of the requirements in the current part 820
  Retains the scope and a number of the definitions from the current part 820
Incorporates by reference ISO 13485:2016
  Minimal called out provisions to ensure consistency with other applicable FDA requirements
    Includes definitions, clarifying concepts, and requirements
Includes conforming edits to Part 4 (cGMPs for combination products)
  Does not impact the CGMP requirements for combination products

## Slide 9: Structure of the Proposed QMSR
820.1 Scope.
820.3 Definitions.
820.7 Incorporation by reference.
820.10 Requirements for a quality management system.
820.15 Clarification of concepts.
820.35 Control of records.
820.45 Device labeling and packaging controls.

## Slide 10: Definitions (§820.3)
Withdraw
Establish
Retain
Act----> FD&C Act
Process Validation-----> Validation of processes
Management with executive responsibility----> Top Management
Finished device
HCT/P
Design validation
Remanufacturer
Nonconformity
Verification
Component

## Slide 11: Definitions (cont.) (§820.3)
Retain with modification
Rework: removing the reference to device master record (DMR)
Clarify/Supersede
Medical device---> Device
Labelling---> Labeling
Add
Customer
Manufacturer
Process agent
Product: retaining, but adding ‘service’ to the definition

## Slide 12: Linkages/ Applicable Regulatory Requirements (§820.10)
Document a QMS
Requirement to comply with other linked/applicable requirements:
  21 CFR Part 830: Unique Device Identification Requirements (Clause 7.5.8)
  21 CFR Part 821: Traceability Requirements, if applicable (Clause 7.5.9)
  21 CFR Part 803: Reporting to Regulatory Authorities (Clause 8.2.3)
  21 CFR Part 806, Advisory Notices (Clauses 7.2.3, 8.2.3, 8.3.3)
Applicability of Design and Development activities
Traceability
  Add a requirement to ensure that devices that support or sustain life, comply with the traceability requirements, in addition to just implantable devices as outlined in Clause 7.5.9.2

## Slide 13: Clarification of Concepts (§820.15)
Organization
Clarify the term to also include the meaning of the term manufacturer as defined in the proposed §820.3
Safety and Performance
Where safety and performance is used, it shall be construed to mean the same as “safety and effectiveness”
Validation of processes
Clarify the term as used in ISO 13485 to refer to “process validation” as defined in the current Part 820

## Slide 14: Control of Records (§820.35)
Signature and date requirements for records
Information required by 21 CFR Part 803, complaint and servicing activities
Documentation required to meet Unique Device Identification (UDI) requirements of 21 CFR Part 830
Confidentiality of records FDA receives
*Must meet these requirements in addition to those in Clause 4.2.5*

## Slide 15: Controls for Device Labeling and Packaging (§820.45)
Proposes to retain requirements from QS reg as ISO 13485 fails to provide additional requirements
ISO 13485 lacks specific requirements to address labeling inspection activities
Intended to strengthen controls for labeling and packaging operations
*Must meet these requirements in addition to those in Clause 7.5.1(e)*

## Slide 16: Proposed QMSR Key Considerations
Does not modify which establishments or products are subject to part 820
Incorporates the 2016 version of ISO 13485
  Any future changes to the standard would need to be evaluated
Proposes a transition period
ISO 13485: 2016 standard available in the ANSI Incorporated by Reference (IBR) Portal
  https://ibr.ansi.org/
FDA will retain our inspectional authority
  FDA inspections will not result in the issuance of certificates of conformance to ISO 13485:2016
  Manufacturers with a certificate of conformance to ISO 13485:2016 are not exempt from FDA inspections
  FDA will not require ISO 13485 certificates

## Slide 17: FDA Implementation Activities
Updating technology systems
Training of personnel
Replacing current inspection model, Quality System Inspection Technique (QSIT)
Revise [and/or develop] relevant regulations, policies, procedures and other documents impacted by this rulemaking
  Compliance Program Guidance Manual (CPGM)
  Guidance Documents
Education and Communication

## Slide 18: Summary
FDA issued a proposed rule to amend Part 820
The proposed rule will incorporate by reference ISO13485:2016
The change will further harmonize global regulatory requirements for QMS
The benefits of this change include simplified and streamlined regulatory requirements and less burden while maintaining a similar level of assurance in a firm’s quality management system