## Slide 1: Linda Moon, Global Quality & Accreditation Manager, BSI Ajith John, Global QMS Technical & Regulatory Reviewer , BSI
MDSAP Assessment Program

## Slide 2: MDSAP Assessment Program – AO Perspective

## Slide 3
Audit Cycle
Initial Audit
Stage 1/Stage 2
Surveillance #1 Audit
Surveillance #2 Audit
Recertification
Special audits
Extension to Scope
Verification visit
Nonconformity close-out
MDSAP Unannounced Audit

## Slide 4
NC Grading
Indirect
Vs
Direct
First occurrence
Vs
Repeat nonconformance
Absence of a process and procedure
Release of nonconforming product
+1
+1
QMS Impact
Escalation Criteria
 | First | Repeat
Indirect | 1 | 2
Direct | 3 | 4
Highest grade is a 5
Repeat = Same subclause X.X.X
Refer GHTF/SG3/N19:2012 for more details and examples

## Slide 5
Post-audit activities
MDSAP report is approved and a report ‘package’ is assembled
100% of MDSAP reports are reviewed by team
Auditors complete audit report
MDSAP team provides regulators with package
Updates are made as needed
Within 45 or 90 days
  Independent certificate review takes places before any MDSAP certificate is issued

## Slide 6
Post-audit activities
D0 | Audit end date (closing meeting)
D0 + 5 W-days | AO informs RA if 1or more Gr 5 NC or more than 2 Gr 4, or public health threat, or fraudulent activity or counterfeit product (MDSAP 5-day Notice)
D0 + 15 C-days | Mfr. provide for each NC a remediation plan (Investigation result of NC and its cause, planned correction and planned corrective action).
D0 + 30 C-days | Mfr. provide for each NC evidence of implementation of the remediation actions taken for any Gr 4 or Gr 5 NC
D0 + 45 C-days | AO to provide complete audit report package if audit meets criteria for a MDSAP 5-day Notice.
D0 + 90 C-days | AO to provide complete audit report package for all other audits.

## Slide 7
MDSAP Unannounced Audits
6-9 months after audit which identified 3+ Grade 4 or 1+ Grade 5 nonconformities
Only required when certificate is already issued (i.e. surveillance or recertification audits)

## Slide 8
MDSAP hybrid audits
Hybrid audits provide much greater flexibility for manufacturers and auditing organisations/notified bodies
Help both parties reduce their carbon footprint and achieve sustainability targets
Regulators are still requiring an on-site component
Particularly useful for manufacturers as experts can join the remote audit stream
MDSAP AU P0036.002 published March 2023
MDSAP - Remote and Hybrid Auditing Pilot Program (fda.gov)
The transmittal is allowed to be implemented immediately – it is a pilot programme for 18 months initially

## Slide 9
Remote/Hybrid audits

## Slide 10
Which manufacturers are eligible for hybrid audits?
QMS primarily with electronic documents and records
Manufacturer with good compliance history
Ability to support an audit via ICT
MDSAP AU P0036.002 defines the tasks to be performed on-site and off-site
List of medical devices where initial or recertification audits are not permitted to be performed using hybrid or remote audit methods

## Slide 11
Question & Answers

## Slide 12
THANK YOU